Gina Litz

GE 330

Sensor Report

pH Sensors and Principles of Operation

How They Work:

Understanding pH measurement

In the process world, pH is an important parameter to be

measured and controlled.

The pH of a solution indicates how acidic or basic (alkaline) it is.

The pH term translates the values of the hydrogen ion concentration

- which ordinarily ranges between about 1 and 10 x -14

gram-equivalents per litre - into numbers between 0 and 14.

On the pH scale a very acidic solution has a low pH value such

as 0, 1, or 2 (which corresponds to a large concentration of

hydrogen ions; 10 x 0, 10 x -1, or 10 x -2 gram-equivalents per litre)

while a very basic solution has a high pH value, such as

12, 13, or 14 which corresponds to a small number of hydrogen

ions (10 x -12, 10 x -13, or 10 x -14 gram-equivalents per litre).

A neutral solution such as water has a pH of approximately 7.

A pH measurement loop is made up of three components, the

pH sensor, which includes a measuring electrode, a reference

electrode, and a temperature sensor; a preamplifier; and an

analyser or transmitter. A pH measurement loop is essentially

a battery where the positive terminal is the measuring electrode

and the negative terminal is the reference electrode. The measuring

electrode, which is sensitive to the hydrogen ion, develops a potential

(voltage) directly related to the hydrogen ion concentration of the

solution. The reference electrode provides a stable potential against

which the measuring electrode can be compared.

Typical pH sensor

When immersed in the solution, the reference electrode potential

does not change with the changing hydrogen ion concentration.

A solution in the reference electrode also makes contact with the

sample solution and the measuring electrode through a junction,

completing the circuit. Output of the measuring electrode changes

with temperature (even though the process remains at a constant pH),

so a temperature sensor is necessary to correct for this change in

output. This is done in the analyser or transmitter software.

The pH sensor components are usually combined into one device

called a combination pH electrode. The measuring electrode is

usually glass and quite fragile. Recent developments have replaced

the glass with more durable solid-state sensors. The preamplifier

is a signal-conditioning device. It takes the high-impedance pH

electrode signal and changes it into alow impedance signal which

the analyser or transmitter can accept. The preamplifier also

strengthens and stabilizes the signal, making it less susceptible

to electrical noise.

The sensor's electrical signal is then displayed. This is commonly

done in a 120/240 V ac-powered analyser or in a 24 V dc loop-powered

transmitter.

Additionally, the analyser or transmitter has a man machine

interface for calibrating the sensor and configuring outputs

and alarms, if pH control is being done.

Keep in mind, application requirements should be carefully

considered when choosing a pH electrode. Accurate pH measurement

and the resulting precise control that it can allow, can go a long

way toward process optimisation and result in increased product

quality and consistency. Accurate, stable pH measurement also

controls and often lowers chemical usage, minimising system

maintenance and expense.

Keeping the system up and running

A system's pH electrodes require periodic maintenance to clean

and calibrate them. The length of time between cleaning and

calibration depends on process conditions and the user's accuracy

and stability expectations. Overtime, electrical properties of the

measuring and reference electrode change. Calibration in

known-value pH solutions called buffers will correct for some

of these changes. Cleaning of the measuring sensor and reference

junction will also help. However, just as batteries have a limited life,

a pH electrode's lifetime is also finite. Even in the "friendliest"

environments, pH electrodes have to be replaced eventually.

Interfacing pH Sensors:

| Ph sensors, interface and calibration |  |
| --- | --- |

The pH electrode is essentially a simple single cell battery. The voltage is directly proportional to the hydrogen ion concentration surrounding the electrode. The pH is the logarithm of the hydrogen ion concentration.

The ideal pH electrode:

- Zero volts output at neutral pH (=7.0)

- Positive voltage in acids, pH<7

- Negative voltages in bases, pH>7

- Total realistic pH range is 0 to 14.

- Generates -59.16 millivolts per pH unit at room temperature (="Nernst potential"). Note that this is a negative slope--higher pH, lower voltage.

- the full scale range is +/-0.414 volts. (+/-0.05916*7), at 25 degC.

- Temperature coefficient of the Nerst potential is -0.001984 mV per °C. That makes the slope -54.2 millivolts per pH unit at 0 degrees Celsius, and -74.04 millivolts per pH unit at 100 degrees Celsius.

But the electrode is practically never ideal:

The above values depend somewhat on the construction of the individual electrode, and its aging. That is why it is necessary to calibrate and standardize the pH monitoring and recording instrument from time to time, depending on the conditions it is subjected to. "Standardize" means to adjust the offset so that the instrument reads zero in neutral (pH 7) solution. "Calibrate" means to trim the slope of the pH/mV response to the the correct value for the electrode at that point in time. The instrument should probably have automatic temperature compensation, to adjust the slope in response to different calibration and working temperatures.

The kicker from the electronic standpoint is that the output impedance of the pH electrode is extremely high. The electrode acts like voltage source, however, there is a 10 to 50 mega ohm resistor in series with the voltage. Any voltmeter that measures the output of a pH electrode has to have extremely high input impedance, 1 tera ohm or more. Even many digital multimeters, which have 10 or 20 mega ohms of input resistance, will load down a pH electrode and give a reading that is much lower than it should be.

It is the glass membrane of the probe that is responsible for the high resistance. It a special glass with tiny "pores" that cannot support much electrical current.

The usual approach to pH electrodes is to amplify and buffer the signal, with an MOS or CMOS input operational amplifier. When properly constructed with attention to the input circuit layout, it can easily achieve the necessary high input resistance.

Here is the circuit of a pHx amplifier:

This amplifier produces 1 volt output in neutral, pH=7, buffer. Adjust RT1 to set this offset. This adjustment could also be used to set the output to 1.0 volts when the pH probe at the input is placed in neutral pH buffer.

The resistor R6 sets the gain. With 221k½ installed at R6, the overall gain will be x2, and the full scale output will be nominally 0.16 to 1.84 volts to cover the 0 to 14 pH range.

This circuit does not have a gain control. The idea is that the main calibration will be done in software. Similarly, although the trimmer RT1 can trim the offset (1 volts output in neutral pH), the idea is to do the recurring calibration in software.

Here is the first cut on how to program it on the Basic Stamp: This assumes that the output of the pH electrode is exactly 0 volts at pH7, and that it has a Nernst slope of exactly 59.16 mV per pH unit:

pHloop:

gosub ADread ' return millivolts, mV, not shown

pH = mV ** 55405      ' stamp's way to multiply times 700/828 (see below)

pH = pH - 845         ' 845 is the 1 volt offset, times 700/828

pH = 700 - pH         ' adjusts to pH 7.00 at 1 volt input, decreasing in acid

debug rep "-"\pH.bit14,dec abs (pH/100),".".dec2 abs pH  ' display with decimal point xx.xx

goto  pHloop

This routine reads the voltage, then converts to pH units, and then prints out the pH result as XX.XX format. The display allows for a negative pH value, but that should never occur except in a super-acid, or unless something is wrong with the instrument.

The factor **55405 is the stamp's way of approximating the fraction (7.00 pH units per 0.828 volt change 700/828), to convert from millivolts to pH. (math: 700/828*65536=55405) The offset -845+700 is applied after the multiplication, because the ** operation does not work correctly on negative numbers. (The offset 845 comes from 1000 millivolts offset * 700/828 pH units per volt).

In the real world, the electrode will not have the ideal slope factor nor the ideal offset. What is more, the slope (Nernst potential) and offset will drift slowly with time as the electrode ages. Also, the pH measurement is temperature dependent. This is predictable, -0.002mV per degree C change in the slope factor.

I practice the pH electrode will have means for interactive calibration. The user first puts the sensor in neutral pH buffer and presses a button, and then puts the probe in either pH4 or pH10 buffer (depending on whether the measurements to come will tend to the acidic or the basic side), and presses a second button. The machine records the readings at those two calibration values and then computes the new slope and offset, which are applied to successive measurements. The temperature is measured too, and the temperature correction is automatically applied. All this can be done using the stamp. More on this later.

Additional considerations apply when measuring pH in real world solutions. Often these solutions contain minerals and other chemical species that affect the temperature dependence of pH in a manner quite different from the theoretical Nernst value of -0.002 mV/degree Celsius. "Cycle Chemistry pH Measurement" is one reference that discusses these issues.

Sending pH sensor readings to the DSP:

The output voltage from the pH sensor must first be converted from mV to between 10 and 10 Volts using 5-B modules.  The 5-B modules shown below will convert the small voltage reading from the pH sensor to between –5 and 5 Volts.

Measurement Computing, Corp. 16 Commerce Blvd. Middleboro, MA 02346 Phone: 508-946-5100 Fax: 508-946-9500 Web: http://www.measurementcomputing.com E-mail: info@measurementcomputing.com ISO-5B30-02 $149.00 +/- 50mV Input, +/-5V Output Analog Isolation Module Features: 1500 Vrms Isolation High Common Mode Rejection Low Cost Specifications: Input Range: +/- 50 mV Full Scale Output Range: +/- 5 V Nonlinearity: +/-0.02% Input Offset: +/-1uV / deg C Output Offset: +/-20uV / deg C Output Resistance: 50 Ohm Bandwidth, -3db: 4 Hz Output Protection: Continuous Short Isolation: 1500 Vrms CMR at 50-60 Hz: 160 db NMR at 50-60 Hz: 90 db Mounting Options: ISO-RACK8 and ISO-RACK16

This analog voltage reading will then be sent to the ADC where it will be converted to a digital input and sent to the DSP.  Then a code would need to be written to convert this new voltage reading to a pH value.  This code would be similar to the code shown above for interfacing.

Most of the pH sensors sold are used for industrial purposes and are already able to convert this value to a pH value, taking into account the change in temperature.  There are different types available based on how they will enter the liquid or flow of liquid.

See data sheets of examples of pH sensors.

References:

http://www.emesystems.com/OWL2face.htm

http://www.sensorland.com/HowPage037.html

http://www.gliint.com/catalog/pdf/data/ds-pc.pdf

http://www.gfsignet.com/products/as2750.html

http://www.measurementcomputing.com