Pressure Sensors (Strain type)

Nathan Hoople

# Measuring Circuits

Bonded resistance strain gage

- Capable of measuring the minute changes in resistance corresponding to strain.

- Usually employ four strain gage elements electrically connected to form a Wheatstone bridge circuit (Figure 2-6).

- A Wheatstone bridge is a divided bridge circuit used for the measurement of static or dynamic electrical resistance.

- Output voltage of the Wheatstone bridge is expressed in millivolts output per volt input.

- The Wheatstone circuit is also well suited for temperature compensation.

Figure 2-6: Wheatstone Bridge Circuit Schematic

- In Figure 2-6, if R1, R2, R3, and R4 are equal, and a voltage, VIN, is applied between points A and C, then the output between points B and D will show no potential difference. However, if R4 is changed to some value, which does not equal R1, R2, and R3, the bridge will become unbalanced and a voltage will exist at the output terminals. In a so-called G-bridge configuration, the variable strain sensor has resistance Rg, while the other arms are fixed value resistors.

- The sensor can occupy one, two, or four arms of the bridge, depending on the application.

- The total strain, or output voltage of the circuit (VOUT) is equivalent to the difference between the voltage drop across R1 and R4, or Rg. This can also be written as:

- The bridge is considered balanced when R1/R2 = Rg/R3 and, therefore, VOUT equals zero.

Pressure Sensors

When a strain gage is used to measure the deflection of an elastic diaphragm or Bourdon tube, it becomes a component in a pressure transducer.

- Used for narrow-span pressure and for differential pressure measurements.

- Measures the displacement of an elastic diaphragm due to a difference in pressure across the diaphragm.

- Detect gauge pressure if the low-pressure port is left open to the atmosphere or differential pressure if connected to two process pressures. If the low-pressure side is a sealed vacuum reference, the transmitter will act as an absolute pressure transmitter.

| Differential pressure tranducers in a variety of ranges and outputs. |  |
| --- | --- |

- Available for pressure ranges as low as 3 inches of water to as high as 200,000 psig (1400 MPa).

- Inaccuracy ranges from 0.1% of span to 0.25% of full scale.

- Additional error sources can be a 0.25% of full-scale drift over six months and a 0.25% full-scale temperature effect per 1000º F.

Figure 1. Shown in this drawing are the major structural elements of a typical high-level pressure sensor.

Pressure Gauge Designs

Two common reasons for gauge (and switch) failure are pipe vibration and water condensation, which in colder climates can freeze and damage the gauge housing. The delicate links, pivots, and pinions of a traditional gauge are sensitive to both condensation and vibration. The life of the filled gauge is longer, not only because it has fewer moving parts, but because its housing is filled with a viscous oil. This oil filling is beneficial not only because it dampens pointer vibration, but also because it leaves no room for humid ambient air to enter. As a result, water cannot condense and accumulate.

- Available gauge features include illuminated dials and digital readouts for better visibility, temperature compensation to correct for ambient temperature variation, differential gauges for differential pressures, and duplex gauges for dual pressure indication on the same dial. Pressure gauges are classified according to their precision, from grade 4A (permissible error of 0.1% of span) to grade D (5% error).

# Protective Accessories

- The most obvious gauge accessory is a shutoff valve between it and the process, which allows blocking while removing or performing maintenance.

- A second valve is often added for one of two reasons: draining of condensate in vapor service (such as steam), or, for higher accuracy applications, to allow calibration against an external pressure source.

- Other accessories include pipe coils or siphons, which in steam service protect the gauge from temperature damage, and snubbers, or pulsation dampeners, which can both absorb pressure shocks and average out pressure fluctuations.

- If freeze protection is needed, the gauge should be heated by steam or electric tracing. Chemical seals protect the gauge from plugging up in viscous or slurry service, and prevent corrosive, noxious or poisonous process materials from reaching the sensor. They also keep the process fluid from freezing or gelling in a dead-ended sensor cavity. The seal protects the gauge by placing a diaphragm between the process and the gauge. The cavity between the gauge and the diaphragm is filled with a stable, low thermal expansion, low viscosity and non-corrosive fluid.

- For high temperature applications, a sodium-potassium eutectic often is used; at ambient temperatures, a mixture of glycerine and water;

- At low temperatures, ethyl alcohol, toluene, or silicon oil.

Interfacing

Output options:

- Small voltage output (0-20mV range)

- Small current output (0-20mA range)

  - Isn’t affected by noise like a voltage output.

  - Good in high noise environment, or if the output signal needs to travel a significant distance to the receiver.

  - Can be converted to a voltage signal using a

- Amplified voltage output (usually between +/-10V)

Can be easily interfaced using the ADC on the DSP.

Other Links:

http://www.omega.com

http://www.entran.com

http://www.vickers-systems.com/products/sensors_subsystems/index.html

http://www.automationsensors.com/frames/indexPRsel.html

http://www.analog.com/

| Entran® Pressure Sensors |  |
| --- | --- |
| EPE Low Pressure Sensors & Microphones Employing the latest technology of micro-machining, the EPE Sensor has been created for low pressure and acoustic testing. Its super sensitive silicon diaphragm has a broad useful range for both static and dynamic pressures and is available in a variety of mounting styles and shapes. Many of these are fitted with a stainless steel protective screen. For use in water or corrosive fluids, contact Entran for special protective coatings. |

Top of Form

Bottom of Form